# Revi Wearable Troubleshooting Guide

General Guidelines and Recommendations

## Who to contact for help?

Please contact your clinician if you have questions about your health or the Revi Wearable Unit.

## Revi Wearable Unit Support:

For questions regarding the Revi Wearable Unit, additional information, and product manuals, visit our website or contact your clinician or BlueWind Patient Support.

NOTE: If any problem remains unresolved, contact your clinician or BlueWind Patient Support so that your Revi system can be evaluated.

| Situation | Possible Cause | Recommended Action |
| --- | --- | --- |
| The Revi Wearable Unit does not turn On. | Insufficient battery power. | Check the battery status by connecting it to the battery charger. If the battery indicator is blinking RED, charge the unit until the LED light turns GREEN. |
|  | Daily Therapy was completed. (The unit has already been activated for the maximum activation time for today). | Wait until tomorrow, then try again. |
|  | The device is connected to the battery charger. | Disconnect the device from the battery charger, then start again. |
| The Revi Wearable Unit does not turn Off. | Power button was not pressed long enough. | Press and hold the Power button for at least 5 seconds. |

| Situation | Possible Cause | Recommended Action |
| --- | --- | --- |
| Cannot find the correct Position for the Revi Wearable Unit. (Positioning beeps are not sounding; the LED indicator is ORANGE and blinking slowly and/or the Warable makes clicking sounds). | The unit is not positioned in the correct location. | Make sure the unit is placed as shown |
|  |  | Move the unit around slowly until you start hearing the Positioning beeps and frequency of the LED blinking increases. |
|  |  | Move the unit to a higher location. Rotate the unit so that the buttons move from the side of your leg to the front of your leg. |
|  | A metal object, or something causing electromagnetic interference is located close to the Revi Wearable Unit. Refer to Electromagnetic Interference and Metal Objects and Implants of the Patient Therapy Guide. | Move the object away from the unit (alternatively, move your leg away from the source of interference). |
| Cannot find the correct Position for the Revi Wearable Unit. (Positioning beeps are not sounding; the LED indicator is not illuminated). | The Wearable Device is turned Off. | Press the Power button once (short press). The LED light will turn ORANGE. Sound and LED will indicate positioning. |
| Beeping sound audible in the middle of Therapy | Device moved out of position. | Re-position the device, then press the Power button once (short press). |
|  | Electromagnetic interference or metal object near the Wearable Unit. | Move away from the interference or move the object away from unit. |
| The Revi Wearable Unit feels excessively hot. | Technical problem. | Make sure you are not wearing the unit while charging the battery. |
|  |  | Contact your clinician for technical support. |

| Situation | Possible Cause | Recommended Action |
| --- | --- | --- |
| Stimulation sensation is lower than usual. | The unit is not in the correct location. | Follow the positioning instructions in the Patient Therapy Guide or in the Revi Wearable Quick Start Guide. |
|  |  | Choose the most optimal position achievable, then start the Therapy Session. |
|  | Maximum Therapy Session level has not been reached. | Press the button to increase Stimulation level. |
|  | Your unit may need reprogramming with new parameters. | Contact your clinician for a new Therapy setup. |
| Revi Wearable Unit does not respond when pressing the + or - buttons. | Maximal / Minimal level has been reached (indicated by three consecutive short beeps). | You may need to contact your clinician for a new Therapy setup. |
|  | For patients for whom Stimulation parameters have been programmed to have On-Off cycles: The + or - button has been pressed during time the unit was in the Off phase. | Press the or button only during time the unit is in the On phase. |
|  | The + button has been pressed during the ramp-up stage (LED blinking GREEN). | Wait for the ramp-up stage to complete, then press the + button (LED light is steady GREEN). |
| Revi Wearable Unit does not start or continue requested action (the LED light is RED). | Any system error. | Retry performing the requested action. If the LED light remains RED, contact your clinician or BlueWind patient support. |
| Feel a change in stimulation when standing up, walking, or changing positions. | Moving the body can affect how the stimulation feels, even though the stimulation level does not change. | Contact your clinician if the stimulation feels unpleasant while sitting. The stimulation may need to be adjusted. |

## This troubleshooting guide is not intended to replace the warnings, cautions and operating instructions provided in the Revi Patient Therapy Guide.

NOTE: Your clinician is your primary source for all questions and requests regarding your Revi System.

As an additional resource, you may contact the manufacturer, BlueWind Medical Ltd. at Tel: (888) 715-2080 or at info@bluewindmedical.com

Dec 2025. G02-CLU-0140 Rev. 2

©2025 BlueWind Medical Ltd.